## Slide 1: From motor cortex to cursor: how Neuralink turns spikes into actions
DRAFT · NEURALINK N1 DATA PATH
Eight stages take the data from neurons firing in the hand-knob region to clicks, typing and a robot arm. On the way it shrinks from about 205 Mbps to roughly 1 Mbps, and then to a few intents per second.
Notes: Compiled 17 Oct 2026 from Neuralink's own updates, the 2019 JMIR paper, a Neuralink patent and the compression challenge. Every fact is tagged with its source in pipeline_data.json.

## Slide 2: Four stages run inside the skull, four on the phone or laptop
THE WHOLE PATH
INSIDE THE N1 IMPLANT
01
Neural source
Motor cortex fires when a hand movement is intended
→
02
Threads
1,024 electrodes on 64 polymer threads
→
03
Amplify + digitise
10-bit ADC, ~20 kHz per channel: ≈205 Mbps
→
04
Spike detection
>200× reduction to spike events
EXTERNAL DEVICE (NEURALINK APP)
05
Radio + power
Through the scalp at ≈1 Mbps; inductive charging
→
06
Decoding
Pretrained Mamba encoder + small heads
→
07
Post-processing
Mode switcher maps intents to clicks, scroll, drag
→
08
Outputs
Cursor, 40 WPM typing, robot arm
Notes: The split matters: everything inside the implant is constrained by power (<10 mW) and radio bandwidth (~1 Mbps). Everything outside can run large ML models.

## Slide 3: The implant listens to the part of the brain that plans hand movement
STAGES 1–2 · SIGNAL SOURCE AND SENSING
1,024
electrodes on 64 threads, each thinner than a hair
~3,000
planned for the next generation
Target found before surgery with fMRI while the patient attempts hand movement (precentral gyrus).
These neurons still fire when a movement is intended, even with a severed spinal cord.
One implant in one hemisphere gives decodable signals for both hands.
The R1 robot inserts the threads. In 2019 it placed up to 6 threads a minute with 87.1% success.
2019 threads: polyimide with gold traces, 5–50 µm wide, 4–6 µm thick.
Notes: Sources: PRIME Study Progress Update (2024-04-12), Two Years of Telepathy (2026-01-28), JMIR 2019 paper.

## Slide 4: Each channel is amplified, filtered and digitised on the implant
STAGE 3 · ANALOG FRONT END AND ADC
Parameter | 2019 research ASIC
Gain | 42.9–59.4 dB
Bandwidth | 3 Hz – 27 kHz
Input noise | 5.9 µVrms
ADC | 10-bit @ 19.3 kHz
Power | 5.2 µW per channel
256 channels per chip. Production N1 chip specs aren't published.
Raw data the N1 produces
204.8 Mbps
1,024 ch × 20,000 samples/s × 10 bit
= 204,800,000 bit/s
Notes: The 204.8 Mbps figure uses the compression challenge's numbers (1,024 channels, ~20 kHz, 10-bit). The table is the 2019 chip, the only one with published specs.

## Slide 5: The bottleneck: the radio carries about 1/200th of what is recorded
STAGE 4 · ON-IMPLANT REDUCTION
Chart: Data rate, log scale (Mbps)
| Raw broadband | Best lossless (~4.4×) | Radio budget
Data rate (Mbps) | 204.8 | 46.5 | 1
Must run in real time: under 1 ms, on less than 10 mW.
Lossless compression stops around 4–5×, so the implant sends spike events and features, not waveforms.
Patent method: filter, fit a model, compare to thresholds, then package each spike with a timestamp.
After threads pulled out in 2024, Neuralink retuned this step to pick up population activity.
Notes: This stage is the least documented. The exact on-implant algorithm isn't public; it's inferred from the patent, the compression challenge (page now offline) and community analysis (Spikey, ~4.4× lossless).

## Slide 6: Spike data leaves the skull by radio, with no wires or connectors
STAGE 5 · WIRELESS LINK AND POWER
1
Data out
Through the scalp to the Neuralink App on a phone or computer, at about 1 Mbps.
2
Power in
Battery charges inductively from the N1 Charger. Nothing passes through the skin.
3
Radio type
Usually described as Bluetooth Low Energy class. Neuralink hasn't confirmed it.
FOR COMPARISON: 2019 RIG
Wired: USB-C to a base station streaming over 10 GbE.
6.5 h/day
average independent use by Feb 2025, which wireless operation made possible

## Slide 7: Pretraining on 50,000 hours of brain data cut calibration time
STAGE 6 · DECODING, THEN AND NOW
BEFORE · PER-USER SUPERVISED
OCT 2026 · PRETRAINED ENCODER
Each user had a decoder for each feature, trained on calibration tasks.
Neural signals drift, so decoders went stale.
About 10 min a day recalibrating (55 min a week).
Self-supervised encoder trained on >50,000 h of unlabeled data.
30 s of labels now matches 3.5 min of old calibration.
Decoders last more than a week; one still worked after 20 months.
Transferred to another participant with 99% of weights frozen.
Notes: Source: Neuralink, Pretraining on 50,000 Hours of Unlabeled Brain Data (2026-10-01). The cross-participant result: a P9-trained model beat P2's own decoder.

## Slide 8: Spikes go in as tokens and intents come out of small heads
STAGE 6 · INSIDE THE ENCODER
Spike tokens
Each spike is a token with a learned per-channel embedding
→
Perceiver input
Compresses spike tokens into a fixed set of latents
→
Mamba backbone
State-space model, fully causal and low-latency
→
Decoder heads
Cursor velocity, click, finger/key, 7+ arm DOF
TRAINING OBJECTIVE
Spatially masked auto-Poisson regression. See half the channels, predict spike counts on the other half. No labels needed, so free-use hours become training data.
SCALE
Participant 1 alone: more than 9,000 hours, about 22.4 billion spikes. Direct target-position decoding ("cursor teleportation") exceeded 10 BPS in early tests.

## Slide 9: Decoded intents become ordinary input events
STAGES 7–8 · POST-PROCESSING AND OUTPUTS
POST-PROCESSING
Earlier versions used smoothing, low-velocity suppression and click filtering. Each added lag.
Pretrained decoders are stable enough that several were dropped.
A mode switcher maps intents to left, right and middle click, scroll, zoom, pan and drag.
Computer + phone
Chess, Civ VI, Mario Kart, Fusion 360, Illustrator
40 WPM
Ten-finger "mind keyboard" typing
Robot arm
CONVOY trial: feeding and gesturing
Next
Speech ~140 WPM (VOICE), vision (Blindsight)

## Slide 10: Cursor control now beats a typical able-bodied mouse user
HOW THE OUTPUT IS MEASURED
Chart: Webgrid bits per second
| P1, first session (early 2024) | P1, May 2024 | Previous record | P15, Oct 2026 (pretrained)
Webgrid BPS | 4.6 | 8 | 10.39 | 11.32
8–10 BPS
able-bodied person using a mouse
21 participants as of Jan 2026, with zero serious device-related adverse events reported. All figures are company-reported.
BPS = log2(N_click · grid² − 1) · max(correct − incorrect, 0) / 60

## Slide 11: Most of the path is documented, but stage 4 is inferred
CONFIDENCE AND GAPS
Stage | Basis | Confidence
1, 2, 5, 6, 7, 8 | Neuralink posts, 2019 peer-reviewed paper | primary
3 · front end | 2019 research chip; production N1 specs unpublished | partial
4 · on-implant | Patent, compression challenge (offline), community analysis | secondary
5 · radio type | Press describes it as BLE-class; not confirmed | secondary
Performance | Company-reported, not independently checked | unverified

## Slide 12: Where each fact comes from
SOURCES
Musk E, Neuralink. An Integrated Brain-Machine Interface Platform With Thousands of Channels. J Med Internet Res 2019;21(10):e16194. doi:10.2196/16194
Neuralink. PRIME Study Progress Update (2024-04-12)
Neuralink. PRIME Update: User Experience (2024-05-08)
Neuralink. PRIME Update: Second Participant (2024-08-21)
Neuralink. A Year of Telepathy (2025-02-05)
Neuralink. Two Years of Telepathy (2026-01-28)
Neuralink. Pretraining on 50,000 Hours of Unlabeled Brain Data (2026-10-01)
Neuralink. Technology page
Neuralink Compression Challenge README (now offline)
US 2021/0012909 A1. Real-time neural spike detection
Spikey, community lossless compression analysis
Notes: Full URLs are in WORKFLOW.md and pipeline_data.json.